## Sheet1
Pre-Bid Queries Responses |  |  |  | 21.06.2025
Open Tender Notification For Supply of Testing Instruments for TPCODL (Tender No - TPCODL/P&S/2025-26/1000002871)
Sr. No. | Detailed Reference to TPCODL Technical Document. Please specify Document No / Clause No / Page No | Description as per Bid Document | Remarks - Query / Clarification | TPCODL Response
1 | 2 | 3 | 4 | 5
Items | Transformer Turns Ratio Meter
1 | Clause No. 1 | This specification covers the technical requirements of design, manufacture, testing at manufacturer's works, packing, forwarding, supply and unloading at store/site of Transformer turns ratio meter complete with all accessories for efficient and trouble free operation. | Please Clarify which type of transformer do you want to test ? | Power & Distribution transformers
2 | Clause No. 3 | The service conditions shall be as follows: 6. Average number of thunderstorm days per annum (isokeraunic level) 70 7. Average number of rainy days per annum 120 8. Average annual rainfall 150cm 9. Earthquakes of an intensity in horizontal direction - equivalent to seismic acceleration of 0.3g 10. Earthquakes of an intensity in vertical direction - equivalent to seismic acceleration of 0.15g (g being acceleration due to gravity) 11 .Wind velocity: 300 km/hr, 200 km/hr and 160 km/hr. environmentally, some of the regions, where the work will take place includes coastal areas, subject to high relative humidity, which can give rise to condensation. Onshore winds will frequently be salt laden. On occasions, the combination of salt and condensation may create pollution conditions for outdoor insulators. Some places are in heavily industrial polluted areas. Therefore, Outdoor material and equipment shall be designed and protected for use in exposed, heavily polluted, salty, corrosive and humid coastal atmosphere The design of equipment and accessories shall be suitable to withstand seismic forces corresponding to an acceleration of 0.1 g. | Point No. 6 to 11 is not applicable for this product so please remove these points. | Can be considered without these conditions but Kit should be complied to Safety/EMC/Vibration standards
3 | Clause No. 4 | 1) Turns Ratio Accuracy ±0.20% (0.8 to 4,000) ±0.25% (4,001 to 10,000) For Excitation Current Values no greater than preset value. | 1-2000 ratio is sufficient for Distribution & Power Transformer. So please amend this as Turns Ratio Accuracy ±0.20% 1-2000 ratio. | Low ratio of 0.8 is required for testing of LVRT
4 | Clause No. 4 | 2) Temperature Range Operating: -20º C to 55º C (-5º F to 130º F) Storage: -40º C to 60º C (-55º F to 140º F) | As per clause no. 3 amend this as Operating: 0º C to 50º C Storage: 0º C to 50º C | As per Specification
5 | Clause No. 4 | 4) Protective Devices High voltage side shorting relay, transient voltage suppressors and gas surge voltage protectors | Here we understand that High voltage side shorting relay means Over Current Relay is it Correct ? | This is for protection of KIt & transformer during testing. Kindly mention protection scheme of proposed kit
6 | Clause No. 4 | 6)Phase Angle Measurement (three phase) - 179.9 ° to + 180 ° Max. error ± 0.05 ° | AS per IEC 60076-1 & ANSI/IEEE C57.12.90 Phase angle measurement is not required so please remove this . | As per Specification
7 | Clause No. 4 | b) Mag.Current Measurement 1 mA – 2000 mA ± 1 mA | please amend this as 1 mA – 500 mA ± 1 mA as per clause 4.9. | As per Specification
8 | Clause No. 4 | c) Current (rms) Accuracy 5% rdg ±0.5 mA | please amend this ±0.1% rdg | Can be Changed
9 | Clause No. 4 | 7)Display Full Graphics LCD module, adjustable back-lighting, wide temperature range, 128 x 64 dots (21 characters by 8 lines) | please amend this 20 x 4 Jumbo Alphanumeric LCD Display | As per Specification
10 | Clause No. 4 | 8)Safety/EMC/Vibration Meets the requirements of IEC-1010-1, CE and ASTM D999.75 | please amend this for IEC 61010-1 | Equivalent IEC standard as per specification can be accepted
11 | Clause No. 4 | 10)Excitation Voltage with Selectable steps upto 100 V AC, 50Hz | please amend this as Excitation voltage should be ≥ 40V which is sufficient. | As per Specification
12 | Clause No. 4 | 11) Excitation Current range 0 to 2000 mA, 4-digit resolution | please amend this as 0 to 200mA, 4-digit Resolution | As per Specification
13 | Clause No. 4 | 12) Auxiliary Supply 230V ± 15% | AC 230V ± 10% | Can be accepted
14 | Clause No. 4 | 14)Weight as furnished by the bidder. Suitable to carry by hand less than 10 KG | 15 to 17kg | As per Specification
15 | Clause No. 4 | 15) Storage Result storage capacity of 500 No’s of results along with Tag of Transformer details. | please amend this 200 Tap Results we can store | As per Specification
16 | Clause No. 4 | 16) Connectivity: able to connect laptop via Bluetooth, LAN & USB cable for data downloading | please amend this as Connectivity: able to connect laptop via Bluetooth/ LAN/USB as suitable for communication between computer & Instrument. | Can be accepted
17 | Clause No. 4 | 18) Battery life operation Up to 12 hours of field operation | up to 1.5Hrs | As per Specification
18 | Clause No. 4 | 20) Test results assessment Kit should automatically asses test results and provide pass or fail feature with editable tolerance limit | please amend this Clause | As per Specification
19 | Clause No. 5 | The Bidder shall furnish the type test certificates for the tests which is required as per the corresponding standards. All the tests shall be conducted at CPRI/ERDA or any other International Laboratory as per the relevant standards. Type tests shall have been conducted in certified Test laboratories during the period not exceeding 5 years from the date of opening the bid. In the event of any discrepancy in the test reports, i.e. any test report not acceptable or any/all type tests (including additional type tests, if any) not carried out, same shall be carried out without any cost implication to TPSODL, Berhampur. | There is no specific standard for Turns ratio meter so please mention standard & list of Tests which are required. | This cause refers to all the type test required for Safety, EMC, Vibration.. as mentioned in type test requirements clause no.
Items | CT Analyser for Protection Current Transformers
20 | Clause No. 05 of 4. GENERAL TECHNICAL PARTICULARS |  | Voltage display is not needed, however if you mean Knee point voltage then it is ok. | Voltage display is required to display during knee point curve and also whereever necessary
21 | Clause No. 11 of 4. GENERAL TECHNICAL PARTICULARS |  | Make current as 15 Apeak | 15A Peak also can be accepted
22 | Clause No. 25 of 4. GENERAL TECHNICAL PARTICULARS |  | Remanence & residual Magnetism Please confirm if Remanence and residual Magnetism shall be measured and its value displayed. | Optional